## Slide 1: Light Calibration System: Preliminary Characterization
Thorsten Lux

## Slide 2
Ar
Air
Light box
X2
X2
WA105: SETUP-SUMMARY
 CIEAMT | IFAE 

## Slide 3: Implementation (Conceptual)
Kaputschinsky PCB
Reference sensor
Reference Sensor
Position of LED adjustable
LED
Feedthrough
Fiber
Stray light
central reference sensor
6 Kaputschinsky PCBs around
each Kaputschinsky with light cavity to guide light to reference sensor
material: either 3D printed plastic or aluminium

## Slide 5: First PCB Test
Photodiode
LED
7 PCBs assembled
tested directly in front of photodiode
2 of 5 showed strange behaviour (e.g. too much light)
5 additional will be assembled
+/- 30% power output between PCBs
PCB | 1 | 2 | 3 | 4 | 5 | 6 | 7
Max. Pulse Energy [pJ] @ 19 V | 30.3 | X | 20.2 | X | 23.9 | 19.5 | 31.3
Max. Power@ 100 Hz [nW] @ 19 V | 3 |  | 2 |  | 2.4 | 2 | 3.1

## Slide 6: Test Setup
Blackbox Wall
Structure
Kaput
Photodiode
Beaglebone
Controller
Optical fiber
all pieces as in final setup
using 1 m fiber, 1 mm diameter reduces power to 1/3
~ 1 nW @ 100 Hz

## Slide 8: Fiber Diameter Dependence
smaller diameter => less power
follows well cross section area
using 50 um instead of 1000 um gives reduction factor of about 400

## Slide 9: Positioning Effects
1
2
3
6

## Slide 10
same PCB and fiber used at all 6 positions
difference between positions less than +/- 10%

## Slide 11
same PCB, position and fiber used at all 5 rotation positions
difference between positions less than +/- 10%

## Slide 12
LED voltage increased from 0 to 19 V
Output power at end of fiber measured with powermeter
pulse rate varied between 100 and 10100 Hz
for plot all power values normalized to 1 kHz
some fluctuations at the beginning due to too short time between setting voltage and start reading => might contribute to previous variations
all curves lay perfectly over the other => system can be operated between 100 Hz and 10 kHz
Frequency Sweep
V
Output power norm to 1 kHz
1 s
2.5 s
5 s

## Slide 13
same PCB, position and fiber used at different gaps
Reduction factor from 0 to 7.5 mm gap about 3-4

## Slide 14: Optical Attenuators
Blackbox Wall
Structure
Kaput
Photodiode
Beaglebone
Controller
1000 um fiber
“1000-standard sleeve-400” 70% of only 400 um fiber
sleeves with different gaps available
400 um fiber
Mating sleeves with gap

## Slide 15
using 2 fibers (1000 and 400 um) with gap mating sleeve gives another reduction factor of 4-5
using a standard SMA + gap + 400 um fiber gives a reduction factor of 20-25 => increasing hole in SMA does not seem to change this much

## Slide 16: Alternative Reduction Idea
replace 100 pF by smaller value

## Slide 17: Requiered Power/ Pulse Energy
From Clara’s talk in September:
up to ~ 1 nW@100Hz@19V with 1 mm diameter fiber => 10 pJ per light pulse
up to 20 pW@100Hz requested => 0.2 pJ per light pulse
large variety of options to match PCB output to required power at flange input
simplest would be to use 400 um fiber instead of 1000 um fiber => 2 pJ per light pulse which could be reduced to 2/3 by positioning of PCB and finally limiting the LED voltage to up to 13 V instead of 19 V
Open question for us still: makes 100 Hz limitation sense in detector with ~10 kHz of cosmics and ~200 kHz of Ar39?

## Slide 18: Status
20 PCBs were ordered, 7 assembled, 5 working and 5 additional are currently assembled
PCBs were tested and the whole system characterized preliminary
Differences between PCBs much larger than positioning issues
Various reduction methods tested to match the output power to the one needed to get single photons
Most of pieces at IFAE => final fibers to be ordered
From this side on track for tests end of February
Reference sensor/DAQ/... => Joan’s talk