## Slide 1: Light Calibration System
Thorsten Lux

## Slide 2: Overall Conceptual Design
Air
X2
Cryostat
Black Box
2 feedthroughs CF40, each with 3 optical feedthroughs
6 SMA optical fibers from feedthrough to black box
black box containing light source
black box on top of cryostat
Inside:
Clara’s talk 2 weeks ago

## Slide 3: Components
Commercial box from Thorlabs: 525 mm x 375 mm x 300 mm (L x W x H)
Kaputschinsky LED drivers as light source (https://arxiv.org/ftp/physics/papers/0410/0410281.pdf)
Provides fast pulses of few ns
Rate: 1 kHz (possible to increase to around 10 kHz)
2 voltages needed:
  DC between 0 and -20 V for intensity
  pulsed to trigger light pulse
Default LED: NSPB300B (465 nm) (more later)
BeagleBoard (BeagleBoard.org)
Easy control via ethernet e.g. with Python
8 outputs to provide pulses for Kaputschinsky and also TTL signal for trigger
7 ADC channels e.g. to read temperature sensor

## Slide 4: Components
SMA-SMA fiber from Thorlabs (1000 um diameter): M59L
SMA-SMA feedthrough from Thorlabs: ADASMA
Optical fiber on one side – LED directly on other side
Reference sensor still open point
Powermeter seem not to work (see later)
Either SiPM or PMT
In contact with Cayetano to check if spare readout channel is available to digitize reference sensor
Alternative: use ADC of BeagleBoard

## Slide 5: Black Box
1
2
3
4
6
5
BeagleBoard
Reference Sensor
Kaputschinsky PCB
Optical fiber
6 Kaputschinsky LED driver
All 6 controlled by 1 BeagleBoard
LEDs can be switched on individually, in groups or all together
each LED coupled to an individual SMA optical fiber
3 fibers going to one CF40 flange
One DC voltage for all 6 Kaputschinsky

## Slide 6: LED Coupling to Optical Fiber
Black box
Kaputschinsky with blue LED
Optical fiber
Stray light used for reference sensor
PMT/SiPM

## Slide 7: Implementation (Conceptual)
Kaputschinsky PCB
Reference sensor
Reference Sensor
Position of LED adjustable
LED
Feedthrough
Fiber
Stray light
central reference sensor
6 Kaputschinsky PCBs around
each Kaputschinsky with light cavity to guide light to reference sensor
material: either 3D printed plastic or aluminium

## Slide 8: Proof-of-Principle Tests
LED
Powermeter (measuring light reaching feedthrough)
PMT
reflector
Reflective tube
Optical fiber

## Slide 9: Comparison PMT and Powermeter
Chart: LED Bias Voltage [V]
| 5 | 6.5 | 7.5 | 9 | 10 | 12.1 | 14 | 16 | 18
Powermeter | 0 | 0.00028329 | 0.0031161 | 0.031445 | 0.084419 | 0.25581 | 0.47365 | 0.72663 | 1.0006
Reference PMT | 0 | 0 | 0.0041667 | 0.029167 | 0.075 | 0.20833 | 0.41667 | 0.66667 | 1
Signal from PMT compared with signal from powermeter
normalized to largest signal
similar shape (amplitude of PMT read by eye from oscilloscope => large error)
=> “stray light” detected by reference light will allow to estimate injected light to PMTs

## Slide 10: Testing various LEDs
Chart: LED Bias Voltage
| 5 | 6.5 | 7.5 | 9 | 10 | 12.1 | 14 | 16 | 18 | 19.5
Blue LED (465 nm) | -0.02 | -0.02 | 0 | 0.22 | 0.58 | 2.08 | 4.4 | 7.83 | 12.12 | 16.07
Green LED (525 nm) | -0.02 | -0.01 | 0.09 | 1.09 | 2.96 | 9.01 | 16.7 | 25.63 | 35.3 | 42.4
Testing various LEDs: 465 nm, 525 nm, 3 of 430 nm
the 3 LEDs of 430 nm do not/ almost not work with Kaputschinsky
Trying to find other 430 nm LEDs
Green LED more power than blue LED
15 nW (blue) correspond to 3.5E10 photons of 465 nm per pulse (1 kHz) reaching the feedthrough on top of the cryostat

## Slide 11: Time Structure of Light Pulse
Old blue LED
New blue LED
Old green LED
New green LED
Green LED with longer tail than blue LED
New green LED with much longer tail than old green LED (same model)
No difference for blue LEDs
Pulse width: 15 ns for blue LED

## Slide 12: CAD Drawings from Juli
drawings from Juli already done
small changes to light cavity necessary
if Joan agrees on PCB size, soon being printed full system
central part flexible so that can be adjusted to SiPM or PMT choice for reference sensor

## Slide 13: CAD Drawings from Juli

## Slide 14: Summary
The light calibration system be based on 6 Kaputschinsky LED driver
Blue LED (465 nm) default choice currently
BeagleBoard will allow to control drivers (individually, groups, all together)
1 reference sensor: SiPM or PMT
Either standard ADC to digitize signal or built-in ADC
Simplified test of the principle performed successfully
Most of the commercial parts ordered
Production drawings in process for light cavity and LED driver PCB
MWPC: CO2 gas will arrive this week, mixture in 3 weeks
DNA Setup: will try to use EL PMTs to avoid noise ripple